## Notes
The calculations in this workbook are based on those in
BS 1042 Part 1 Section 1 : 1981
The calculations are valid for Diameter Ratioes between 0.2 and 0.6 (0.2 < ß < 0.6)
Uncertainty (tolerance) is based on worse-case scenario
The results cannot be trusted if your system is out of tolerance with any part of the standard
Populate the yellow boxes
You can access the calculations and hidden sheets by unprotecting sheets and the workbook as appropriate

## Results Sheet
|  |  | (min) | (max)
q(v) @ 10" H2O = | 207.9035 | CFM | 201.2246 | 213.0398
q(v) @ 25" H2O = | 328.7243 | CFM | 318.164 | 336.8454
q(v) @ 28" H2O = | 347.8891 | CFM | 336.7131 | 356.4837
Diameter Ratio = | 0.59558 |  | 0.59104 | 0.60089

## Inputs - Dimensions
BS 1042 Part 1 Section 1 : 1981
Temperature at Measuring Station (Tm) | 15.5 | °C | Tm(min) = | 15 | Tm(max) = | 16
Bore Diameter in Region of Orifice Plate (see 6.1.5) | 1 | 2 | 3
i - v | 65.15 | 65.04 | 65.23
ii - vi | 64.77 | 65.14 | 65.07
iii - vii | 64.73 | 64.9 | 65.15
iiii - viii | 65.16 | 65.15 | 65.3
D = | 65.0658 | Error
D(min) = | 64.73 | 0.5161% | The system is outside the scope of BS1042 if either of these values is greater than 0.3%
D(max) = | 65.3 | 0.8806%
Orifice Diameter (see 7.1.7.2) | 1 | 2 | 3
i - v | 38.7 | 38.62 | 38.68
ii - vi | 38.7 | 38.75 | 38.74
iii - vii | 38.85 | 38.87 | 38.87
iiii - viii | 38.74 | 38.76 | 38.74
d = | 38.7517 | Error
d(min) = | 38.62 | 0.3398% | The system is outside the scope of BS1042 if either of these values is greater than 0.05%
d(max) = | 38.87 | 0.6473%
Orifice Thickness (see Figure 2, 7.1.4.1 & 7.1.4.2) | 1 | 2 | 3
i | 0.73 | 0.77 | 0.78
ii | 0.76 | 0.73 | 0.8
iii | 0.52 | 0.58 | 0.55
iiii | 0.85 | 0.81 | 0.66
v | 0.68 | 0.76 | 0.76
vi | 0.65 | 0.62 | 0.63
vii | 0.68 | 0.58 | 0.69
viii | 0.6 | 0.66 | 0.72
e = | 0.69042 | Error | Max Allowable Error
e(min) = | 0.52 | 0.17042 | 0.065066 |  | The system is outside the scope of BS1042 if either of these values is greater than the max allowable error (0.001D)
e(max) = | 0.85 | 0.15958
Orifice Plate Thickness (see Figure 2, 7.1.4.3 & 7.1.4.4) | 1 | 2 | 3
i | 3.18 | 3.13 | 3.18
ii | 2.96 | 3 | 3.03
iii | 2.96 | 2.91 | 2.93
iiii | 3.23 | 3.15 | 3.08
v | 3.25 | 3.25 | 3.22
vi | 2.95 | 2.97 | 2.95
vii | 3.01 | 3.03 | 3.03
viii | 3.1 | 3.09 | 3.08
E = | 3.0696 | Allowed
E(min) = | 2.91 | 0.69042 | The system is outside the scope of BS1042 if E(min) is less than the minimum thickness or E(max) is greater than the maximum allowable thickness
E(max) = | 3.25 | 3.2533
Distance to Upstream Tap (see Figure 3 & 7.2.3.4) | 1 | 2 | 3
i | 63.48 | 63.68 | 63.44
l1 = | 63.5333 | Allowed
l1(min) = | 63.44 | 58.5592 | The system is outside the scope of BS1042 if l1(min) is less 0.9D or l1(max) is greater than 1.1D
l1(max) = | 63.68 | 71.5724
Distance to Downstream Tap (see Figure 3 & 7.2.3.4) | 1 | 2 | 3
i | 33.39 | 33.47 | 33.53
l2 = | 33.4633 | Allowed
l2(min) = | 33.39 | 31.2316 | The system is outside the scope of BS1042 if l2(min) is less 0.48D or l2(max) is greater than 0.52D
l2(max) = | 33.53 | 33.8342

## Inputs - Non-dimensional
BS 1042 Part 1 Section 1 : 1981
Pressure Difference Across Test Orifice (mm H2O) | 69.9 | mm H2O
| 70.1 | mm H2O
| 70.3 | mm H2O
Test P = | 70.1 | mm H2O
Test P(min) = | 69.9 | mm H2O
Test P(max) = | 70.3 | mm H2O
Pressure Difference Across Measurement Orifice (mm H2O) | 285 | mm H2O
| 305 | mm H2O
| 310 | mm H2O
Delta P = | 300 | mm H2O
Delta P(min) = | 285 | mm H2O
Delta P(max) = | 310 | mm H2O
Ambient Pressure (mBar) | 1013.5 | mBar
| 1013.7 | mBar
| 1013.9 | mBar
Amb P = | 1013.7 | mBar
Amb P(min) = | 1013.5 | mBar
Amb P(max) = | 1013.9 | mBar
Pressure Difference Between Downstream Tap & Atmosphere | 110 | mm H2O
| 100 | mm H2O
| 120 | mm H2O
Absolute Downstream Tap Pressure P2 = | 10226.8633 | mm H2O
P2(min) = | 10214.8238 | mm H2O
P2(max) = | 10238.9027 | mm H2O
Upstream Tap Pressure P1 = | 10636.8633 | mm H2O
P1(min) = | 10644.8238 | mm H2O
P1(max) = | 10623.9027 | mm H2O
Pressure Ratio (See 2.2.4 & 5.3.3)
P2/P1 = | 0.96145 | The system is outside the scope of BS1042 if any of these values is less than 0.75
Pressure Ratio (min) = | 0.9596
Pressure Ratio (max) = | 0.96376
Ambient Temperature (°C) | 20.5 | °C
| 20.4 | °C
| 20.9 | °C
Amb T = | 20.6 | °C
Amb T(min) = | 20.4 | °C
Amb T(max) = | 20.9 | °C
Temperature at the Upstream Tap (°C) | 35.3 | °C
(Measured 5D Downstream from Orifice Plate) | 34.6 | °C
(see 3.4.2) | 36.2 | °C
T1 = | 35.3667 | °C
T1(min) = | 34.6 | °C
T1(max) = | 36.2 | °C

## Inputs - Internet
| min | nominal | max
Temperature (T1) = | 34.6 | 35.3667 | 36.2 | °C
Pressure (P1) = | 1043.9006 | 1043.12 | 1041.849 | mBar
Humidity = | 49 | 50 | 51 | %
Dew Point = | 22 | 23 | 24 | °C | http://www.dpcalc.org/
Air Density @ Upstream Tapping Temperature = | 1.1014 | 1.099 | 1.0977 | kg/cu m | https://www.gribble.org/cycling/air_density.html
Temperature (Amb 1) = | 20.4 | 20.6 | 20.9 | °C
Pressure (Amb P) = | 1013.5 | 1013.7 | 1013.9 | mBar
Humidity = | 49 | 50 | 51 | %
Dew Point = | 9 | 10 | 11 | °C | http://www.dpcalc.org/
Air Density @ Upstream Tapping Temperature = | 1.1976 | 1.1967 | 1.1953 | kg/cu m | https://www.gribble.org/cycling/air_density.html
Coefficient of Themal Expansion = | 90 | 100 | 110 | (micro metres per metre per Kelvin)
https://www.engineeringtoolbox.com/linear-expansion-coefficients-d_95.html
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05 | N-s/m2 | https://www.lmnoeng.com/Flow/GasViscosity.php
Isentropic Exponent (k) = | 1.4 | 1.401 | 1.401 | https://www.engineeringtoolbox.com/specific-heat-ratio-d_602.html

## Dimensions and Readings (hidden)
Dimensions and Readings |  |  |  |  | Common Calculations
| min | nominal | max |  |  | min | nominal | max
Temperature at Measuring Station (Tm) | 15 | 15.5 | 16 | °C | Diameter Ratio (unadjusted for temp) (ß) | 0.59142 | 0.59558 | 0.60049
Bore Diameter (unadjusted for temp) (D) | 64.73 | 65.0658 | 65.3 | mm | Relative Upstream Tap Ratio (unadjusted for temp) (L1) | 0.97152 | 0.97645 | 0.98378
Orifice Diameter (unadjusted for temp) (d) | 38.62 | 38.7517 | 38.87 | mm | Relative Downstream Tap Ratio (unadjusted for temp) (L2) | 0.51133 | 0.5143 | 0.518
Orifice Thickness (e) | 0.52 | 0.69042 | 0.85 | mm | Upstream Pressure (P1) | 10049.8238 | 10036.8633 | 10028.9027 | mm H2O
Orifice Plate Thickness (E) | 2.91 | 3.0696 | 3.25 | mm | Bore Diameter (adjusted for temperature) (D) | 64.8384 | 65.1951 | 65.4523 | mm
Distance to Upstream Tap (unadjusted for temp) (l1) | 63.44 | 63.5333 | 63.68 | mm | Orifice Diameter (adjusted for temperature) (d) | 38.6846 | 38.8287 | 38.9606 | mm
Distance to Downstream Tap (unadjusted for temp) (l2) | 33.39 | 33.4633 | 33.53 | mm | Distance to Upstream Tap (adjusted for temp) (l1) | 63.5462 | 63.6596 | 63.8285 | mm
Pressure Difference Across Test Orifice (Test Delta P) | 69.9 | 70.1 | 70.3 | mm H2O | Distance to Downstream Tap (adjusted for temp) (l2) | 33.4459 | 33.5298 | 33.6082 | mm
Pressure Difference Across Measuring Orifice (Delta P) | 285 | 300 | 310 | mm H2O | Diameter Ratio (adjusted for temp) (ß) | 0.59104 | 0.59558 | 0.60089
Ambient Pressure (Amb P) | 1013.5 | 1013.7 | 1013.9 | mBar | Relative Upstream Tap Ratio (adjusted for temp) (L1) | 0.97088 | 0.97645 | 0.98443
Ambient Temperature (Amb T) | 20.4 | 20.6 | 20.9 | °C | Relative Downstream Tap Ratio (adjusted for temp) (L2) | 0.511 | 0.5143 | 0.51834
Temperature at Upstream Tap (T1) | 34.6 | 35.3667 | 36.2 | °C | Expansion Factor (€) | 0.99083 | 0.99031 | 0.98995
Air Density at Upstream Tap (Rho 1) | 1.1014 | 1.099 | 1.0977 | kg/cu m | Approach Velocity Factor (E) | 1.0672 | 1.0695 | 1.0723
Air Density at Test Orifice (Rho Test) | 1.1976 | 1.1967 | 1.1953 | kg/cu m
Coefficient of Thermal Expansion | 90 | 100 | 110 | µm/mK | Volume Flow Rate [ q(v) ] | 105.5608 | 109.2204 | 112.0783 | CFM
Dynamic Viscosity (µ) | 1.9084e-05 | 1.9123e-05 | 1.9162e-05 | N-s/m2
Isentropic Exponent (k) | 1.4 | 1.401 | 1.401

## Re (1st iteration) (hidden)
Re(1) = | 1000000

## C (1st iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
D = | 64.8384 | 65.1951 | 65.4523
ß = | 0.59663 | 0.59558 | 0.59525
l1 = | 63.5462 | 63.6596 | 63.8285
l2 = | 33.4459 | 33.5298 | 33.6082
L1 = | 0.98007 | 0.97645 | 0.97519
L2 = | 0.51584 | 0.5143 | 0.51348
Re(1) = | 1000000 | 1000000 | 1000000
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.010547 | 0.010508 | 0.010496
Term 3 = | -0.0029544 | -0.0029128 | -0.0029002
Term 4 = | 0.00079738 | 0.00079386 | 0.00079278
Term 5 = | 0.0055462 | 0.005481 | 0.0054604
Term 6 = | -0.003692 | -0.0036615 | -0.0036497
C(1) = | 0.60614 | 0.60611 | 0.6061

## Mass Flow Rate (1st iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 285 | 300 | 310 | mm H2O
Delta P = | 2794.8183 | 2941.914 | 3039.9778 | N/sq m
q(m) (1) = | 0.059111 | 0.061129 | 0.062665 | kg/s

## Re (2nd iteration) (hidden)
| min | nominal | max
q(m) (1) = | 0.059111 | 0.061129 | 0.062665
D = | 64.8384 | 65.1951 | 65.4523
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05
Re(2) = | 60823.1002 | 62428.882 | 63617.3187

## C (2nd iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
D = | 64.8384 | 65.1951 | 65.4523
ß = | 0.59663 | 0.59558 | 0.59525
l1 = | 63.5462 | 63.6596 | 63.8285
l2 = | 33.4459 | 33.5298 | 33.6082
L1 = | 0.98007 | 0.97645 | 0.97519
L2 = | 0.51584 | 0.5143 | 0.51348
Re(2) = | 60823.1002 | 62428.882 | 63617.3187
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.010547 | 0.010508 | 0.010496
Term 3 = | -0.0029544 | -0.0029128 | -0.0029002
Term 4 = | 0.0065105 | 0.0063563 | 0.0062585
Term 5 = | 0.0055462 | 0.005481 | 0.0054604
Term 6 = | -0.003692 | -0.0036615 | -0.0036497
C(2) = | 0.61186 | 0.61167 | 0.61157

## Mass Flow Rate (2nd iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 285 | 300 | 310 | mm H2O
Delta P = | 2794.8183 | 2941.914 | 3039.9778 | N/sq m
q(m) (2) = | 0.059668 | 0.06169 | 0.06323 | kg/s

## Re (3rd iteration) (hidden)
| min | nominal | max
q(m) (2) = | 0.059668 | 0.06169 | 0.06323
D = | 64.8384 | 65.1951 | 65.4523
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05
Re(3) = | 61396.3774 | 63001.8079 | 64191.0105

## C (3rd iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
D = | 64.8384 | 65.1951 | 65.4523
ß = | 0.59663 | 0.59558 | 0.59525
l1 = | 63.5462 | 63.6596 | 63.8285
l2 = | 33.4459 | 33.5298 | 33.6082
L1 = | 0.98007 | 0.97645 | 0.97519
L2 = | 0.51584 | 0.5143 | 0.51348
Re(3) = | 61396.3774 | 63001.8079 | 64191.0105
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.010547 | 0.010508 | 0.010496
Term 3 = | -0.0029544 | -0.0029128 | -0.0029002
Term 4 = | 0.0064648 | 0.0063129 | 0.0062165
Term 5 = | 0.0055462 | 0.005481 | 0.0054604
Term 6 = | -0.003692 | -0.0036615 | -0.0036497
C(3) = | 0.61181 | 0.61163 | 0.61152

## Mass Flow Rate (3rd iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 285 | 300 | 310 | mm H2O
Delta P = | 2794.8183 | 2941.914 | 3039.9778 | N/sq m
q(m) (3) = | 0.059663 | 0.061685 | 0.063225 | kg/s

## Re (4th iteration) (hidden)
| min | nominal | max
q(m) (3) = | 0.059663 | 0.061685 | 0.063225
D = | 64.8384 | 65.1951 | 65.4523
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05
Re(4) = | 61391.7971 | 62997.3375 | 64186.6024

## C (4th iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
D = | 64.8384 | 65.1951 | 65.4523
ß = | 0.59663 | 0.59558 | 0.59525
l1 = | 63.5462 | 63.6596 | 63.8285
l2 = | 33.4459 | 33.5298 | 33.6082
L1 = | 0.98007 | 0.97645 | 0.97519
L2 = | 0.51584 | 0.5143 | 0.51348
Re(4) = | 61391.7971 | 62997.3375 | 64186.6024
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.010547 | 0.010508 | 0.010496
Term 3 = | -0.0029544 | -0.0029128 | -0.0029002
Term 4 = | 0.0064652 | 0.0063132 | 0.0062168
Term 5 = | 0.0055462 | 0.005481 | 0.0054604
Term 6 = | -0.003692 | -0.0036615 | -0.0036497
C(4) = | 0.61181 | 0.61163 | 0.61152

## Mass Flow Rate (4th iteration) (hidden)
| min | nominal | max
d = | 38.6846 | 38.8287 | 38.9606
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 285 | 300 | 310 | mm H2O
Delta P = | 2794.8183 | 2941.914 | 3039.9778 | N/sq m
q(m) (4) = | 0.059663 | 0.061685 | 0.063225 | kg/s

## Final Calculations Sheet (hidden)
| min | nominal | max
Air Density = | 1.1976 | 1.1967 | 1.1953 | kg/cu m
Test Orifice Depression = | 69.9 | 70.1 | 70.3 | mm H2O
Test Orifice Depression = | 2.752 | 2.7598 | 2.7677 | in H2O
q(v) = | 0.049819 | 0.051546 | 0.052895 | cu m/s
q(v) = | 105.5608 | 109.2204 | 112.0783 | CFM
q(v) @ 10" H2O = | 201.2246 | 207.9035 | 213.0398 | CFM
q(v) @ 25" H2O = | 318.164 | 328.7243 | 336.8454 | CFM
q(v) @ 28" H2O = | 336.7131 | 347.8891 | 356.4837 | CFM