## Notes
The calculations in this workbook are based on those in
BS 1042 Part 1 Section 1 : 1981
The calculations are valid for Diameter Ratioes between 0.2 and 0.6 (0.2 < ß < 0.6)
Uncertainty (tolerance) is based on worse-case scenario
The results cannot be trusted if your system is out of tolerance with any part of the standard
Populate the yellow boxes
You can access the calculations and hidden sheets by unprotecting sheets and the workbook as appropriate

## Results Sheet
|  |  | (min) | (max)
q(v) @ 10" H2O = | 135.368 | CFM | 135.5899 | 135.2725
q(v) @ 25" H2O = | 214.0356 | CFM | 214.3864 | 213.8846
q(v) @ 28" H2O = | 226.514 | CFM | 226.8852 | 226.3542
Diameter Ratio = | 0.49767 |  | 0.49762 | 0.49772

## Inputs - Dimensions
BS 1042 Part 1 Section 1 : 1981
Temperature at Measuring Station (Tm) | 15 | °C | Tm(min) = | 14.5 | Tm(max) = | 15.5
Bore Diameter in Region of Orifice Plate (see 6.1.5) | 1 | 2 | 3
i - v | 64.3 | 64.3 | 64.3
ii - vi | 64.3 | 64.3 | 64.3
iii - vii | 64.3 | 64.3 | 64.3
iiii - viii | 64.3 | 64.3 | 64.3
D = | 64.3 | Error
D(min) = | 64.3 | 0% | The system is outside the scope of BS1042 if either of these values is greater than 0.3%
D(max) = | 64.3 | 0%
Orifice Diameter (see 7.1.7.2) | 1 | 2 | 3
i - v | 32 | 32 | 32
ii - vi | 32 | 32 | 32
iii - vii | 32 | 32 | 32
iiii - viii | 32 | 32 | 32
d = | 32 | Error
d(min) = | 32 | 0% | The system is outside the scope of BS1042 if either of these values is greater than 0.05%
d(max) = | 32 | 0%
Orifice Thickness (see Figure 2, 7.1.4.1 & 7.1.4.2) | 1 | 2 | 3
i | 0.73 | 0.77 | 0.78
ii | 0.76 | 0.73 | 0.8
iii | 0.52 | 0.58 | 0.55
iiii | 0.85 | 0.81 | 0.66
v | 0.68 | 0.76 | 0.76
vi | 0.65 | 0.62 | 0.63
vii | 0.68 | 0.58 | 0.69
viii | 0.6 | 0.66 | 0.72
e = | 0.69042 | Error | Max Allowable Error
e(min) = | 0.52 | 0.17042 | 0.0643 |  | The system is outside the scope of BS1042 if either of these values is greater than the max allowable error (0.001D)
e(max) = | 0.85 | 0.15958
Orifice Plate Thickness (see Figure 2, 7.1.4.3 & 7.1.4.4) | 1 | 2 | 3
i | 3.18 | 3.13 | 3.18
ii | 2.96 | 3 | 3.03
iii | 2.96 | 2.91 | 2.93
iiii | 3.23 | 3.15 | 3.08
v | 3.25 | 3.25 | 3.22
vi | 2.95 | 2.97 | 2.95
vii | 3.01 | 3.03 | 3.03
viii | 3.1 | 3.09 | 3.08
E = | 3.0696 | Allowed
E(min) = | 2.91 | 0.69042 | The system is outside the scope of BS1042 if E(min) is less than the minimum thickness or E(max) is greater than the maximum allowable thickness
E(max) = | 3.25 | 3.215
Distance to Upstream Tap (see Figure 3 & 7.2.3.4) | 1 | 2 | 3
i | 25.4 | 25.4 | 25.4
l1 = | 25.4 | Allowed
l1(min) = | 25.4 | 57.87 | The system is outside the scope of BS1042 if l1(min) is less 0.9D or l1(max) is greater than 1.1D
l1(max) = | 25.4 | 70.73
Distance to Downstream Tap (see Figure 3 & 7.2.3.4) | 1 | 2 | 3
i | 25.4 | 25.4 | 25.4
l2 = | 25.4 | Allowed
l2(min) = | 25.4 | 30.864 | The system is outside the scope of BS1042 if l2(min) is less 0.48D or l2(max) is greater than 0.52D
l2(max) = | 25.4 | 33.436

## Inputs - Non-dimensional
BS 1042 Part 1 Section 1 : 1981
Pressure Difference Across Test Orifice (mm H2O) | 69.9 | mm H2O
| 70.1 | mm H2O
| 70.3 | mm H2O
Test P = | 70.1 | mm H2O
Test P(min) = | 69.9 | mm H2O
Test P(max) = | 70.3 | mm H2O
Pressure Difference Across Measurement Orifice (mm H2O) | 300 | mm H2O
| 300 | mm H2O
| 300 | mm H2O
Delta P = | 300 | mm H2O
Delta P(min) = | 300 | mm H2O
Delta P(max) = | 300 | mm H2O
Ambient Pressure (mBar) | 1013.25 | mBar
| 1013.25 | mBar
| 1013.25 | mBar
Amb P = | 1013.25 | mBar
Amb P(min) = | 1013.25 | mBar
Amb P(max) = | 1013.25 | mBar
Pressure Difference Between Downstream Tap & Atmosphere | 0 | mm H2O
| 0 | mm H2O
| 0 | mm H2O
Absolute Downstream Tap Pressure P2 = | 10332.2745 | mm H2O
P2(min) = | 10332.2745 | mm H2O
P2(max) = | 10332.2745 | mm H2O
Upstream Tap Pressure P1 = | 10632.2745 | mm H2O
P1(min) = | 10632.2745 | mm H2O
P1(max) = | 10632.2745 | mm H2O
Pressure Ratio (See 2.2.4 & 5.3.3)
P2/P1 = | 0.97178 | The system is outside the scope of BS1042 if any of these values is less than 0.75
Pressure Ratio (min) = | 0.97178
Pressure Ratio (max) = | 0.97178
Ambient Temperature (°C) | 15 | °C
| 15 | °C
| 15 | °C
Amb T = | 15 | °C
Amb T(min) = | 15 | °C
Amb T(max) = | 15 | °C
Temperature at the Upstream Tap (°C) | 15 | °C
(Measured 5D Downstream from Orifice Plate) | 15 | °C
(see 3.4.2) | 15 | °C
T1 = | 15 | °C
T1(min) = | 15 | °C
T1(max) = | 15 | °C

## Inputs - Internet
| min | nominal | max
Temperature (T1) = | 15 | 15 | 15 | °C
Pressure (P1) = | 1042.67 | 1042.67 | 1042.67 | mBar
Humidity = | 49 | 50 | 51 | %
Dew Point = | 22 | 23 | 24 | °C | http://www.dpcalc.org/
Air Density @ Upstream Tapping Temperature = | 1.1014 | 1.099 | 1.0977 | kg/cu m | https://www.gribble.org/cycling/air_density.html
Temperature (Amb 1) = | 15 | 15 | 15 | °C
Pressure (Amb P) = | 1013.25 | 1013.25 | 1013.25 | mBar
Humidity = | 49 | 50 | 51 | %
Dew Point = | 9 | 10 | 11 | °C | http://www.dpcalc.org/
Air Density @ Upstream Tapping Temperature = | 1.1976 | 1.1967 | 1.1953 | kg/cu m | https://www.gribble.org/cycling/air_density.html
Coefficient of Themal Expansion = | 90 | 100 | 110 | (micro metres per metre per Kelvin)
https://www.engineeringtoolbox.com/linear-expansion-coefficients-d_95.html
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05 | N-s/m2 | https://www.lmnoeng.com/Flow/GasViscosity.php
Isentropic Exponent (k) = | 1.4 | 1.401 | 1.401 | https://www.engineeringtoolbox.com/specific-heat-ratio-d_602.html

## Dimensions and Readings
Dimensions and Readings |  |  |  |  | Common Calculations
| min | nominal | max |  |  | min | nominal | max
Temperature at Measuring Station (Tm) | 14.5 | 15 | 15.5 | °C | Diameter Ratio (unadjusted for temp) (ß) | 0.49767 | 0.49767 | 0.49767
Bore Diameter (unadjusted for temp) (D) | 64.3 | 64.3 | 64.3 | mm | Relative Upstream Tap Ratio (unadjusted for temp) (L1) | 0.39502 | 0.39502 | 0.39502
Orifice Diameter (unadjusted for temp) (d) | 32 | 32 | 32 | mm | Relative Downstream Tap Ratio (unadjusted for temp) (L2) | 0.39502 | 0.39502 | 0.39502
Orifice Thickness (e) | 0.52 | 0.69042 | 0.85 | mm | Upstream Pressure (P1) | 10032.2745 | 10032.2745 | 10032.2745 | mm H2O
Orifice Plate Thickness (E) | 2.91 | 3.0696 | 3.25 | mm | Bore Diameter (adjusted for temperature) (D) | 64.2971 | 64.3 | 64.3035 | mm
Distance to Upstream Tap (unadjusted for temp) (l1) | 25.4 | 25.4 | 25.4 | mm | Orifice Diameter (adjusted for temperature) (d) | 31.9986 | 32 | 32.0018 | mm
Distance to Downstream Tap (unadjusted for temp) (l2) | 25.4 | 25.4 | 25.4 | mm | Distance to Upstream Tap (adjusted for temp) (l1) | 25.3989 | 25.4 | 25.4014 | mm
Pressure Difference Across Test Orifice (Test Delta P) | 69.9 | 70.1 | 70.3 | mm H2O | Distance to Downstream Tap (adjusted for temp) (l2) | 25.3989 | 25.4 | 25.4014 | mm
Pressure Difference Across Measuring Orifice (Delta P) | 300 | 300 | 300 | mm H2O | Diameter Ratio (adjusted for temp) (ß) | 0.49762 | 0.49767 | 0.49772
Ambient Pressure (Amb P) | 1013.25 | 1013.25 | 1013.25 | mBar | Relative Upstream Tap Ratio (adjusted for temp) (L1) | 0.39498 | 0.39502 | 0.39506
Ambient Temperature (Amb T) | 15 | 15 | 15 | °C | Relative Downstream Tap Ratio (adjusted for temp) (L2) | 0.39498 | 0.39502 | 0.39506
Temperature at Upstream Tap (T1) | 15 | 15 | 15 | °C | Expansion Factor (€) | 0.99078 | 0.99079 | 0.99079
Air Density at Upstream Tap (Rho 1) | 1.1014 | 1.099 | 1.0977 | kg/cu m | Approach Velocity Factor (E) | 1.0321 | 1.0322 | 1.0322
Air Density at Test Orifice (Rho Test) | 1.1976 | 1.1967 | 1.1953 | kg/cu m
Coefficient of Thermal Expansion | 90 | 100 | 110 | µm/mK | Volume Flow Rate [ q(v) ] | 71.1294 | 71.1145 | 71.1656 | CFM
Dynamic Viscosity (µ) | 1.9084e-05 | 1.9123e-05 | 1.9162e-05 | N-s/m2
Isentropic Exponent (k) | 1.4 | 1.401 | 1.401

## Re (1st iteration)
Re(1) = | 1000000

## C (1st iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
D = | 64.2971 | 64.3 | 64.3035
ß = | 0.49767 | 0.49767 | 0.49767
l1 = | 25.3989 | 25.4 | 25.4014
l2 = | 25.3989 | 25.4 | 25.4014
L1 = | 0.39502 | 0.39502 | 0.39502
L2 = | 0.39502 | 0.39502 | 0.39502
Re(1) = | 1000000 | 1000000 | 1000000
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.0072065 | 0.0072065 | 0.0072065
Term 3 = | -0.00069236 | -0.00069236 | -0.00069236
Term 4 = | 0.00050669 | 0.00050669 | 0.00050669
Term 5 = | 0.0010068 | 0.0010068 | 0.0010068
Term 6 = | -0.0016409 | -0.0016409 | -0.0016409
C(1) = | 0.60229 | 0.60229 | 0.60229

## Mass Flow Rate (1st iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 300 | 300 | 300 | mm H2O
Delta P = | 2941.914 | 2941.914 | 2941.914 | N/sq m
q(m) (1) = | 0.039873 | 0.039834 | 0.039815 | kg/s

## Re (2nd iteration)
| min | nominal | max
q(m) (1) = | 0.039873 | 0.039834 | 0.039815
D = | 64.2971 | 64.3 | 64.3035
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05
Re(2) = | 41373.2163 | 41247.1212 | 41142.3395

## C (2nd iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
D = | 64.2971 | 64.3 | 64.3035
ß = | 0.49767 | 0.49767 | 0.49767
l1 = | 25.3989 | 25.4 | 25.4014
l2 = | 25.3989 | 25.4 | 25.4014
L1 = | 0.39502 | 0.39502 | 0.39502
L2 = | 0.39502 | 0.39502 | 0.39502
Re(2) = | 41373.2163 | 41247.1212 | 41142.3395
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.0072065 | 0.0072065 | 0.0072065
Term 3 = | -0.00069236 | -0.00069236 | -0.00069236
Term 4 = | 0.0055234 | 0.0055361 | 0.0055466
Term 5 = | 0.0010068 | 0.0010068 | 0.0010068
Term 6 = | -0.0016409 | -0.0016409 | -0.0016409
C(2) = | 0.6073 | 0.60732 | 0.60733

## Mass Flow Rate (2nd iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 300 | 300 | 300 | mm H2O
Delta P = | 2941.914 | 2941.914 | 2941.914 | N/sq m
q(m) (2) = | 0.040205 | 0.040166 | 0.040148 | kg/s

## Re (3rd iteration)
| min | nominal | max
q(m) (2) = | 0.040205 | 0.040166 | 0.040148
D = | 64.2971 | 64.3 | 64.3035
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05
Re(3) = | 41717.8313 | 41591.5528 | 41486.6182

## C (3rd iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
D = | 64.2971 | 64.3 | 64.3035
ß = | 0.49767 | 0.49767 | 0.49767
l1 = | 25.3989 | 25.4 | 25.4014
l2 = | 25.3989 | 25.4 | 25.4014
L1 = | 0.39502 | 0.39502 | 0.39502
L2 = | 0.39502 | 0.39502 | 0.39502
Re(3) = | 41717.8313 | 41591.5528 | 41486.6182
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.0072065 | 0.0072065 | 0.0072065
Term 3 = | -0.00069236 | -0.00069236 | -0.00069236
Term 4 = | 0.0054891 | 0.0055016 | 0.0055121
Term 5 = | 0.0010068 | 0.0010068 | 0.0010068
Term 6 = | -0.0016409 | -0.0016409 | -0.0016409
C(3) = | 0.60727 | 0.60728 | 0.60729

## Mass Flow Rate (3rd iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 300 | 300 | 300 | mm H2O
Delta P = | 2941.914 | 2941.914 | 2941.914 | N/sq m
q(m) (3) = | 0.040203 | 0.040164 | 0.040146 | kg/s

## Re (4th iteration)
| min | nominal | max
q(m) (3) = | 0.040203 | 0.040164 | 0.040146
D = | 64.2971 | 64.3 | 64.3035
Dynamic Viscosity (µ) = | 1.9084e-05 | 1.9123e-05 | 1.9162e-05
Re(4) = | 41715.4782 | 41589.1956 | 41484.2576

## C (4th iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
D = | 64.2971 | 64.3 | 64.3035
ß = | 0.49767 | 0.49767 | 0.49767
l1 = | 25.3989 | 25.4 | 25.4014
l2 = | 25.3989 | 25.4 | 25.4014
L1 = | 0.39502 | 0.39502 | 0.39502
L2 = | 0.39502 | 0.39502 | 0.39502
Re(4) = | 41715.4782 | 41589.1956 | 41484.2576
See BS1042-1.1:1987, 7.3.2.1
Term 1 = | 0.5959 | 0.5959 | 0.5959
Term 2 = | 0.0072065 | 0.0072065 | 0.0072065
Term 3 = | -0.00069236 | -0.00069236 | -0.00069236
Term 4 = | 0.0054894 | 0.0055019 | 0.0055123
Term 5 = | 0.0010068 | 0.0010068 | 0.0010068
Term 6 = | -0.0016409 | -0.0016409 | -0.0016409
C(4) = | 0.60727 | 0.60728 | 0.60729

## Mass Flow Rate (4th iteration)
| min | nominal | max
d = | 31.9986 | 32 | 32.0018
Air Density = | 1.1014 | 1.099 | 1.0977
Delta P = | 300 | 300 | 300 | mm H2O
Delta P = | 2941.914 | 2941.914 | 2941.914 | N/sq m
q(m) (4) = | 0.040203 | 0.040164 | 0.040146 | kg/s

## Final Calculations Sheet
| min | nominal | max
Air Density = | 1.1976 | 1.1967 | 1.1953 | kg/cu m
Test Orifice Depression = | 69.9 | 70.1 | 70.3 | mm H2O
Test Orifice Depression = | 2.752 | 2.7598 | 2.7677 | in H2O
q(v) = | 0.033569 | 0.033562 | 0.033586 | cu m/s
q(v) = | 71.1294 | 71.1145 | 71.1656 | CFM
q(v) @ 10" H2O = | 135.5899 | 135.368 | 135.2725 | CFM
q(v) @ 25" H2O = | 214.3864 | 214.0356 | 213.8846 | CFM
q(v) @ 28" H2O = | 226.8852 | 226.514 | 226.3542 | CFM