## Notes
The calculations in this workbook are based on those in
ISO 5167-1:2003(E) & ISO 5167-2:2003(E)
The calculations are valid for Diameter Ratioes between 0.2 and 0.6 (0.2 < ß < 0.6)
Uncertainty (tolerance) is based on worse-case scenario
The results cannot be trusted if your system is out of tolerance with any part of the standard
Populate the yellow boxes | Link to ISO 5167-1:2003(E) | Link to ISO 5167-2:2003(E)
You can access the calculations and hidden sheets by unprotecting sheets and the workbook as appropriate
After you have populated "Inputs - Constants", you shouldn't need to change any of the values unless you change your system

## Results Sheet
|  |  | (min) | (max)
q(v) @ 10" H2O = | 363.4913 | CFM | 367.3769 | 359.9183
q(v) @ 25" H2O = | 574.7303 | CFM | 580.8738 | 569.0807
q(v) @ 28" H2O = | 608.2374 | CFM | 614.7391 | 602.2585
Discharge Coefficient (C) = | 0.61423 |  | 0.61485 | 0.61365
Diameter Ratio = | 0.5938 |  | 0.58247 | 0.60596

## Inputs - Constants
ISO 5167-2:2003(E)
Temperature at Measuring Station (Tm) | 16.4 | °C | Tm(min) = | 15.9 | Tm(max) = | 16.9
Bore Diameter in Region of Orifice Plate (see 6.4.2) | 1 | 2 | 3
i - v | 65.25 | 65.21 | 65.23
ii - vi | 65.22 | 65.06 | 65.25
iii - vii | 66.07 | 66.29 | 66.22
iiii - viii | 64.27 | 64.16 | 64.9
D = | 65.2608 | Error
D(min) = | 64.16 | 1.687% | The system is outside the scope of ISO 5167-2 if either of these values is greater than 0.3%
D(max) = | 66.29 | 3.32%
Orifice Diameter (see 5.1.8) | 1 | 2 | 3
i - v | 38.7 | 38.62 | 38.68
ii - vi | 38.7 | 38.75 | 38.74
iii - vii | 38.85 | 38.87 | 38.87
iiii - viii | 38.74 | 38.76 | 38.74
d = | 38.7517 | Error
d(min) = | 38.62 | 0.3398% | The system is outside the scope of ISO 5167-2 if either of these values is greater than 0.05%
d(max) = | 38.87 | 0.6473%
Orifice Thickness (see Figure 1, 5.1.5.1 & 5.1.5.2) | 1 | 2 | 3
i | 0.73 | 0.77 | 0.78
ii | 0.76 | 0.73 | 0.8
iii | 0.52 | 0.58 | 0.55
iiii | 0.85 | 0.81 | 0.66
v | 0.68 | 0.76 | 0.76
vi | 0.65 | 0.62 | 0.63
vii | 0.68 | 0.58 | 0.69
viii | 0.6 | 0.66 | 0.72
e = | 0.69042 | Error | Max Allowable Error
e(min) = | 0.52 | 0.17042 | 0.065261 |  | The system is outside the scope of ISO 5167-2 if either of these values is greater than the max allowable error (0.001D)
e(max) = | 0.85 | 0.15958
Orifice Plate Thickness (see Figure 1, 5.1.5.3 & 5.1.5.4) | 1 | 2 | 3
i | 3.18 | 3.13 | 3.18
ii | 2.96 | 3 | 3.03
iii | 2.96 | 2.91 | 2.93
iiii | 3.23 | 3.15 | 3.08
v | 3.25 | 3.25 | 3.22
vi | 2.95 | 2.97 | 2.95
vii | 3.01 | 3.03 | 3.03
viii | 3.1 | 3.09 | 3.08
E = | 3.0696 | Allowed
E(min) = | 2.91 | 0.69042 | The system is outside the scope of ISO 5167-2 if E(min) is less than the minimum thickness or E(max) is greater than the maximum allowable thickness
E(max) = | 3.25 | 3.263
Distance to Upstream Tap (see Figure 3 & 5.2.2.2) | 1 | 2 | 3
i | 63.48 | 63.68 | 63.44
l1 = | 63.5333 | Allowed
l1(min) = | 63.44 | 58.7347 | The system is outside the scope of ISO 5167-2 if l1(min) is less 0.9D or l1(max) is greater than 1.1D
l1(max) = | 63.68 | 71.7869
Distance to Downstream Tap (see Figure 3 & 5.2.2.2) | 1 | 2 | 3
i | 33.39 | 33.47 | 33.46
l2 = | 33.44 | Allowed
l2(min) = | 33.39 | 31.3252 | The system is outside the scope of ISO 5167-2 if l2(min) is less 0.48D or l2(max) is greater than 0.52D
l2(max) = | 33.47 | 33.9356
Coefficient of Themal Expansion = | 100 | 100 | 100 | (micro metres per metre per Kelvin)
https://www.engineeringtoolbox.com/linear-expansion-coefficients-d_95.html
Latitude = | 52.8984 | 52.8984 | 52.8984 | degrees | (see Google Earth, or similar)
Elevation (altitude) = | 10 | 10 | 10 | metres | (see Ordnance Survey)

## Inputs - Variables
ISO 5167-1:2003(E) & ISO 5167-2:2003(E)
Pressure Difference Across Test Orifice (mm H2O) | 21 | mm H2O
| 23 | mm H2O
| 22 | mm H2O
Test P = | 22 | mm H2O
Test P(min) = | 21 | mm H2O
Test P(max) = | 23 | mm H2O
Pressure Difference Across Measurement Orifice (mm H2O) | 276 | mm H2O
| 276 | mm H2O
| 275 | mm H2O
Delta P = | 275.6667 | mm H2O
Delta P(min) = | 275 | mm H2O
Delta P(max) = | 276 | mm H2O
Ambient Pressure (mBar) | 1012.6 | mBar
(Measured on sea-level calibrated barometer) | 1012.7 | mBar
| 1012.7 | mBar | Corrected
Amb P = | 1012.6667 | mBar | 1014.4242 | 10344.2483
Amb P(min) = | 1012.6 | mBar | 1014.301 | 10342.9922
Amb P(max) = | 1012.7 | mBar | 1014.5196 | 10345.221
Pressure Difference Between Upstream Tap & Atmosphere | 172 | mm H2O
| 174 | mm H2O
| 173 | mm H2O | Corrected + P1
Absolute Upstream Tap Pressure P1 = | 10517.1538 | mm H2O | 1031.3805 | 10517.1538
P1(min) = | 10514.9922 | mm H2O | 1031.1685 | 10514.9922
P1(max) = | 10519.221 | mm H2O | 1031.5832 | 10519.221
Pressure Difference Between Downstream Tap & Atmosphere | 101 | mm H2O
| 101 | mm H2O
| 100 | mm H2O | Corrected + P2
Absolute Downstream Tap Pressure P2 = | 10444.8205 | mm H2O | 1024.287 | 10444.8205
P2(min) = | 10442.9922 | mm H2O | 1024.1077 | 10442.9922
P2(max) = | 10446.221 | mm H2O | 1024.4243 | 10446.221
Pressure Ratio (See 2.2.4 & 5.3.3)
P2/P1 = | 0.99312 | The system is outside the scope of ISO 5167 if any of these values is less than 0.75
Pressure Ratio (min) = | 0.99315
Pressure Ratio (max) = | 0.99306
Ambient Temperature (°C) | 11.7 | °C
| 12 | °C
| 12.4 | °C
Amb T = | 12.0333 | °C
Amb T(min) = | 11.7 | °C
Amb T(max) = | 12.4 | °C
Temperature at the Upstream Tap (°C) | 23.6 | °C
(Measured 5D Downstream from Orifice Plate) | 24.2 | °C
(see 3.4.2) | 24.7 | °C
T1 = | 24.1667 | °C
T1(min) = | 23.6 | °C
T1(max) = | 24.7 | °C
Ambient Humidity | 82 | %
| 82 | %
| 82 | %
Amb H% = | 82
Amb H%(min) = | 82
Amb H%(max) = | 82

## Inputs - Internet
| min | nominal | max
Temperature (T1) = | 23.6 | 24.1667 | 24.7 | °C
Pressure (P1) = | 1031.1685 | 1031.3805 | 1031.5832 | mBar
Dynamic Viscosity @ Upstream Plane (µ) = | 1.8121e-05 | 1.8121e-05 | 1.8121e-05 | N-s/m2 | http://www.peacesoftware.de/einigewerte/luft_e.html
Temperature (T1) = | 23.6 | 24.1667 | 24.7 | °C
Isentropic Exponent @ Upstream Plane (k) = | 1.401 | 1.401 | 1.401 | https://www.engineeringtoolbox.com/specific-heat-ratio-d_602.html

## Air Pressure (hidden)
Ambient
| min | nominal | max
obs lat | 52.8984 | 52.8984 | 52.8984 | Degrees
obs temp | 11.7 | 12.0333 | 12.4 | °C
obs height | 10 | 10 | 10 | metres
obs barometer | 1012.6 | 1012.6667 | 1012.7 | mBar
lat | 52.8984 | 52.8984 | 52.8984 | Degrees
temp | 53.06 | 53.66 | 54.32 | °F
height | 32.8084 | 32.8084 | 32.8084 | Feet
barometer | 29.9021 | 29.904 | 29.905 | in Hg
lat | 0.92325 | 0.92325 | 0.92325 | Radians
temp | 512.73 | 513.33 | 513.99 | °R
temp1 | 53.06 | 53.66 | 54.32 | °F
latitude correction | 0.019984 | 0.019985 | 0.019986 | in Hg
temperature correction | 0.066187 | 0.067813 | 0.069599 | in Hg
altitude correction | -0.035838 | -0.035796 | -0.03575 | in Hg
correction | 0.050333 | 0.052002 | 0.053835 | in Hg
actual pressure | 1014.301 | 1014.4242 | 1014.5196 | mBar

## Dew Point (hidden)
| Ambient |  |  | Upstream Tap Plane
| min | nominal | max | Min | Nominal | Max
T = | 11.7 | 12.0333 | 12.4 | 23.6 | 24.1667 | 24.7 | °C
Humidity = | 82 | 82 | 82 | 82 | 82 | 82 | %
gamma = | 0.61442 | 0.63654 | 0.66082 | 1.3717 | 1.4062 | 1.4385
T(dewpoint) = | 8.7465 | 9.0725 | 9.4311 | 20.3808 | 20.9345 | 21.4556 | °C

## Air Density (hidden)
| Ambient |  |  | Upstream Tap Plane
| Min | Nominal | Max | Min | Nominal | Max
Ambient Temp (T) = | 11.7 | 12.0333 | 12.4 | 23.6 | 24.1667 | 24.7 | °C
Ambient Temp (Tk) = | 284.85 | 285.1833 | 285.55 | 296.75 | 297.3167 | 297.85 | °K
Pressure (P) = | 1014.301 | 1014.4242 | 1014.5196 | 1031.1685 | 1031.3805 | 1031.5832 | mBar
Dewpoint (Tv) = | 8.7465 | 9.0725 | 9.4311 | 20.3808 | 20.9345 | 21.4556 | °C
c0 = 0.99999683 | 1 | 1 | 1 | 1 | 1 | 1
c1 = -0.90826951 · 10-2 | -0.0090827 | -0.0090827 | -0.0090827 | -0.0090827 | -0.0090827 | -0.0090827
c2 = 0.78736169 · 10-4 | 7.8736e-05 | 7.8736e-05 | 7.8736e-05 | 7.8736e-05 | 7.8736e-05 | 7.8736e-05
c3 = -0.61117958 · 10-6 | -6.1118e-07 | -6.1118e-07 | -6.1118e-07 | -6.1118e-07 | -6.1118e-07 | -6.1118e-07
c4 = 0.43884187 · 10-8 | 4.3884e-09 | 4.3884e-09 | 4.3884e-09 | 4.3884e-09 | 4.3884e-09 | 4.3884e-09
c5 = -0.29883885 · 10-10 | -2.9884e-11 | -2.9884e-11 | -2.9884e-11 | -2.9884e-11 | -2.9884e-11 | -2.9884e-11
c6 = 0.21874425 · 10-12 | 2.1874e-13 | 2.1874e-13 | 2.1874e-13 | 2.1874e-13 | 2.1874e-13 | 2.1874e-13
c7 = -0.17892321 · 10-14 | -1.7892e-15 | -1.7892e-15 | -1.7892e-15 | -1.7892e-15 | -1.7892e-15 | -1.7892e-15
c8 = 0.11112018 · 10-16 | 1.1112e-17 | 1.1112e-17 | 1.1112e-17 | 1.1112e-17 | 1.1112e-17 | 1.1112e-17
c9 = -0.30994571 · 10-19 | -3.0995e-20 | -3.0995e-20 | -3.0995e-20 | -3.0995e-20 | -3.0995e-20 | -3.0995e-20
Pv = c0 + T (c1 + T (c2 + T (c3 + T (c4 + T (c5 + T (c6 + T (c7 + T (c8 + T (c9) ) ) ) ) ) ) ) | 0.92619 | 0.92365 | 0.92086 | 0.84308 | 0.83949 | 0.83614 | mBar
Pv = | 92.6194 | 92.3646 | 92.086 | 84.3081 | 83.9493 | 83.6143 | Pa
Pd = P — Pv | 1013.3748 | 1013.5006 | 1013.5988 | 1030.3254 | 1030.541 | 1030.747 | mBar
= | 101337.4846 | 101350.0582 | 101359.8753 | 103032.5397 | 103054.0972 | 103074.704 | Pa
Rv is 461.4964 | 461.495 | 461.495 | 461.495 | 461.495 | 461.495 | 461.495
Rd is 287.0531 | 287.05 | 287.05 | 287.05 | 287.05 | 287.05 | 287.05
Tk is measured temperature in degrees Kelvin, i.e., measured temperature T (Celsius) + 273.15 | 284.85 | 285.1833 | 285.55 | 296.75 | 297.3167 | 297.85 | °K
Rho (kg/m3) = (Pd / (Rd · Tk)) + (Pv / (Rv · Tk)) | 1.2401 | 1.2388 | 1.2373 | 1.2102 | 1.2081 | 1.2062 | kg/cu m

## Dimensions and Readings (hidden)
Dimensions and Readings |  |  |  |  | Common Calculations
| min | nominal | max |  |  | min | nominal | max
Temperature at Measuring Station (Tm) | 15.9 | 16.4 | 16.9 | °C | Diameter Ratio (unadjusted for temp) (ß) | 0.58259 | 0.5938 | 0.60583
Bore Diameter (unadjusted for temp) (D) | 64.16 | 65.2608 | 66.29 | mm | Relative Upstream Tap Ratio (unadjusted for temp) (L1) | 0.95701 | 0.97353 | 0.99252
Orifice Diameter (unadjusted for temp) (d) | 38.62 | 38.7517 | 38.87 | mm | Relative Downstream Tap Ratio (unadjusted for temp) (L2) | 0.5037 | 0.51241 | 0.52166
Orifice Thickness (e) | 0.52 | 0.69042 | 0.85 | mm | Pressure Ratio (P2/P1) | 0.99315 | 0.99312 | 0.99306
Orifice Plate Thickness (E) | 2.91 | 3.0696 | 3.25 | mm | Bore Diameter (adjusted for temperature) (D) | 64.203 | 65.3115 | 66.3483 | mm
Distance to Upstream Tap (unadjusted for temp) (l1) | 63.44 | 63.5333 | 63.68 | mm | Orifice Diameter (adjusted for temperature) (d) | 38.6459 | 38.7818 | 38.9042 | mm
Distance to Downstream Tap (unadjusted for temp) (l2) | 33.39 | 33.44 | 33.47 | mm | Distance to Upstream Tap (adjusted for temp) (l1) | 63.4825 | 63.5827 | 63.736 | mm
[5 empty] | Distance to Downstream Tap (adjusted for temp) (l2) | 33.4124 | 33.466 | 33.4995 | mm
Ambient Pressure (Amb P) | 1014.301 | 1014.4242 | 1014.5196 | mBar | Diameter Ratio (adjusted for temp) (ß) | 0.58247 | 0.5938 | 0.60596
Ambient Temperature (Amb T) | 11.7 | 12.0333 | 12.4 | °C | Relative Upstream Tap Ratio (adjusted for temp) (L1) | 0.95681 | 0.97353 | 0.99273
Pressure Difference Across Test Orifice (Test Delta P) | 21 | 22 | 23 | mm H2O | Relative Downstream Tap Ratio (adjusted for temp) (L2) | 0.50359 | 0.51241 | 0.52177
Air Density at Test Orifice (Rho Test) | 1.2401 | 1.2388 | 1.2373 | kg/cu m | Expansion Factor (€) (see ISO 5167-2:2003(E), 5.3.2.2) | 0.99808 | 0.99805 | 0.998
[5 empty] | Approach Velocity Factor (E) | 1.0631 | 1.0686 | 1.0751
Upstream Pressure (P1) | 10514.9922 | 10517.1538 | 10519.221 | mm H2O | M2 (see ISO 5167-2:2003(E), 5.3.2.1) | 2.4122 | 2.5229 | 2.6483
Temperature at Upstream Tap (T1) | 23.6 | 24.1667 | 24.7 | °C
Air Density at Upstream Tap (Rho 1) | 1.2102 | 1.2081 | 1.2062 | kg/cu m
[5 empty] | Volume Flow Rate [ q(v) ] | 105.6342 | 106.9765 | 108.3056 | CFM
Pressure Difference Across Measuring Orifice (Delta P) | 275 | 275.6667 | 276 | mm H2O
Downstream Pressure (P2) | 10442.9922 | 10444.8205 | 10446.221 | mm H2O
Coefficient of Thermal Expansion | 100 | 100 | 100 | µm/mK
Dynamic Viscosity (µ) | 1.8121e-05 | 1.8121e-05 | 1.8121e-05 | N-s/m2
Isentropic Exponent (k) | 1.401 | 1.401 | 1.401

## Re (1st iteration) (hidden)
Re(1) = | 1000000

## C (1st iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
D = | 64.203 | 65.3115 | 66.3483
ß = | 0.60193 | 0.5938 | 0.58636
l1 = | 63.4825 | 63.5827 | 63.736
l2 = | 33.4124 | 33.466 | 33.4995
L1 = | 0.98878 | 0.97353 | 0.96063
L2 = | 0.52042 | 0.51241 | 0.5049
Re(1) = | 1000000 | 1000000 | 1000000
A = | 0.027967 | 0.027664 | 0.027386
M2 = | 2.4122 | 2.5229 | 2.6483
See ISO 5167-2:2003(E), 5.3.2.1
Term 1 = | 0.5961 | 0.5961 | 0.5961
Term 2 = | 0.0094566 | 0.0092027 | 0.0089737
Term 3 = | -0.0037225 | -0.0033385 | -0.0030185
Term 4 = | 0.000521 | 0.000521 | 0.000521
Term 5 = | 0.0032109 | 0.0030612 | 0.0029289
Term 6 = | 0.0064603 | 0.0060678 | 0.0057282
Term 7 = | -0.0048844 | -0.0048629 | -0.0048466
Term 8 = | 0.00044354 | 0.00039293 | 0.00033815
C(1) = | 0.60759 | 0.60714 | 0.60672

## Mass Flow Rate (1st iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
Air Density = | 1.2102 | 1.2081 | 1.2062
Delta P = | 275 | 275.6667 | 276 | mm H2O
Delta P = | 2696.7545 | 2703.2921 | 2706.5609 | N/sq m
q(m) (1) = | 0.061092 | 0.06182 | 0.06253 | kg/s

## Re (2nd iteration) (hidden)
| min | nominal | max
q(m) (1) = | 0.061092 | 0.06182 | 0.06253
D = | 64.203 | 65.3115 | 66.3483
Dynamic Viscosity (µ) = | 1.8121e-05 | 1.8121e-05 | 1.8121e-05
Re(2) = | 66859.8753 | 66509.1877 | 66221.0691

## C (2nd iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
D = | 64.203 | 65.3115 | 66.3483
ß = | 0.60193 | 0.5938 | 0.58636
l1 = | 63.4825 | 63.5827 | 63.736
l2 = | 33.4124 | 33.466 | 33.4995
L1 = | 0.98878 | 0.97353 | 0.96063
L2 = | 0.52042 | 0.51241 | 0.5049
Re(2) = | 66859.8753 | 66509.1877 | 66221.0691
A = | 0.24351 | 0.24188 | 0.24029
M2 = | 2.4122 | 2.5229 | 2.6483
See ISO 5167-2:2003(E), 5.3.2.1
Term 1 = | 0.5961 | 0.5961 | 0.5961
Term 2 = | 0.0094566 | 0.0092027 | 0.0089737
Term 3 = | -0.0037225 | -0.0033385 | -0.0030185
Term 4 = | 0.0034612 | 0.0034739 | 0.0034845
Term 5 = | 0.0077463 | 0.0073939 | 0.0070807
Term 6 = | 0.0063067 | 0.0059244 | 0.0055937
Term 7 = | -0.0048844 | -0.0048629 | -0.0048466
Term 8 = | 0.00044354 | 0.00039293 | 0.00033815
C(2) = | 0.61491 | 0.61429 | 0.61371

## Mass Flow Rate (2nd iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
Air Density = | 1.2102 | 1.2081 | 1.2062
Delta P = | 275 | 275.6667 | 276 | mm H2O
Delta P = | 2696.7545 | 2703.2921 | 2706.5609 | N/sq m
q(m) (2) = | 0.061828 | 0.062548 | 0.063249 | kg/s

## Re (3rd iteration) (hidden)
| min | nominal | max
q(m) (2) = | 0.061828 | 0.062548 | 0.063249
D = | 64.203 | 65.3115 | 66.3483
Dynamic Viscosity (µ) = | 1.8121e-05 | 1.8121e-05 | 1.8121e-05
Re(3) = | 67665.59 | 67291.572 | 66982.9799

## C (3rd iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
D = | 64.203 | 65.3115 | 66.3483
ß = | 0.60193 | 0.5938 | 0.58636
l1 = | 63.4825 | 63.5827 | 63.736
l2 = | 33.4124 | 33.466 | 33.4995
L1 = | 0.98878 | 0.97353 | 0.96063
L2 = | 0.52042 | 0.51241 | 0.5049
Re(3) = | 67665.59 | 67291.572 | 66982.9799
A = | 0.24118 | 0.23963 | 0.2381
M2 = | 2.4122 | 2.5229 | 2.6483
See ISO 5167-2:2003(E), 5.3.2.1
Term 1 = | 0.5961 | 0.5961 | 0.5961
Term 2 = | 0.0094566 | 0.0092027 | 0.0089737
Term 3 = | -0.0037225 | -0.0033385 | -0.0030185
Term 4 = | 0.0034323 | 0.0034456 | 0.0034567
Term 5 = | 0.0077129 | 0.0073629 | 0.0070516
Term 6 = | 0.0063084 | 0.0059259 | 0.0055951
Term 7 = | -0.0048844 | -0.0048629 | -0.0048466
Term 8 = | 0.00044354 | 0.00039293 | 0.00033815
C(3) = | 0.61485 | 0.61423 | 0.61365

## Mass Flow Rate (3rd iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
Air Density = | 1.2102 | 1.2081 | 1.2062
Delta P = | 275 | 275.6667 | 276 | mm H2O
Delta P = | 2696.7545 | 2703.2921 | 2706.5609 | N/sq m
q(m) (3) = | 0.061822 | 0.062542 | 0.063244 | kg/s

## Re (4th iteration) (hidden)
| min | nominal | max
q(m) (3) = | 0.061822 | 0.062542 | 0.063244
D = | 64.203 | 65.3115 | 66.3483
Dynamic Viscosity (µ) = | 1.8121e-05 | 1.8121e-05 | 1.8121e-05
Re(4) = | 67658.923 | 67285.2343 | 66976.9269

## C (4th iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
D = | 64.203 | 65.3115 | 66.3483
ß = | 0.60193 | 0.5938 | 0.58636
l1 = | 63.4825 | 63.5827 | 63.736
l2 = | 33.4124 | 33.466 | 33.4995
L1 = | 0.98878 | 0.97353 | 0.96063
L2 = | 0.52042 | 0.51241 | 0.5049
Re(4) = | 67658.923 | 67285.2343 | 66976.9269
A = | 0.2412 | 0.23965 | 0.23812
M2 = | 2.4122 | 2.5229 | 2.6483
See ISO 5167-2:2003(E), 5.3.2.1
Term 1 = | 0.5961 | 0.5961 | 0.5961
Term 2 = | 0.0094566 | 0.0092027 | 0.0089737
Term 3 = | -0.0037225 | -0.0033385 | -0.0030185
Term 4 = | 0.0034325 | 0.0034458 | 0.0034569
Term 5 = | 0.0077132 | 0.0073631 | 0.0070519
Term 6 = | 0.0063083 | 0.0059259 | 0.005595
Term 7 = | -0.0048844 | -0.0048629 | -0.0048466
Term 8 = | 0.00044354 | 0.00039293 | 0.00033815
C(4) = | 0.61485 | 0.61423 | 0.61365

## Mass Flow Rate (4th iteration) (hidden)
| min | nominal | max
d = | 38.6459 | 38.7818 | 38.9042
Air Density = | 1.2102 | 1.2081 | 1.2062
Delta P = | 275 | 275.6667 | 276 | mm H2O
Delta P = | 2696.7545 | 2703.2921 | 2706.5609 | N/sq m
q(m) (4) = | 0.061822 | 0.062542 | 0.063244 | kg/s

## Final Calculations Sheet (hidden)
| min | nominal | max
Air Density = | 1.2401 | 1.2388 | 1.2373 | kg/cu m
Test Orifice Depression = | 21 | 22 | 23 | mm H2O
Test Orifice Depression = | 0.82677 | 0.86614 | 0.90551 | in H2O
q(v) = | 0.049854 | 0.050487 | 0.051115 | cu m/s
q(v) = | 105.6342 | 106.9765 | 108.3056 | CFM
q(v) @ 10" H2O = | 367.3769 | 363.4913 | 359.9183 | CFM
q(v) @ 25" H2O = | 580.8738 | 574.7303 | 569.0807 | CFM
q(v) @ 28" H2O = | 614.7391 | 608.2374 | 602.2585 | CFM